## Slide 1
MPP Basics
Communication with PRISMAPRO
A Pfeiffer Vacuum© application

## Slide 2: MPP: Application for communication between Mass-Spectrometer PrismaPro and PC / PLC
* Copyright notice: Image source "© Siemens AG 2020, All rights reserved".
** wikipedia: OpenPlatformCommunications-UnifiedArchitecture
*** wikipedia: Application Programming Interface
API*** via http
Start/Stop…
Recipe ….
Values/Conditions …
Recipe, Commands…
S7 / OPC-UA**
(MPP: Application for Communication between Mass-Spectrometer PrismaPro and S7-PLC or OPC-UA-Server)

## Slide 3
  Establish a connection between a PrismaPro and a Siemens PLC or an OPC-UA-Server
  Switch On/Off Emission + EM (and a lot more: See Documentation.pdf)
  Read/Write a recipe from/to the PrismaPro (currently 20 Channels/Recipe)
  Start/Stop a RUN with „Lock“ and „Release“ the PrismaPro for other Applications
  Read the measured values after or during each RUN (max. 3001 Datapoints/RUN)
Cyclic reading: Internal + external Pressure (mBar & Pas & Torr) Faults and Warnings State of Emission, EM, Scan, ….. Digital/Analog Inputs + Outputs (Option)
Feature: Adjust a recipe in PVMassSpec and use with the MPP-application
Main topics:

## Slide 4: Installation 1/4:
Copy all necessary files into a folder on the disk (recommended is a folder named MPP in the root folder C:\).

## Slide 5: Installation 2/4:
Press Windows key and type in cmd. The best match should be the Command Prompt application which must be started with administrator rights (Right click → Run as administrator).

## Slide 6: Installation 3/4:
Type in the following command and confirm it with a press of the Enter key: “C:\Windows\Microsoft.NET\Framework\v4.0.30319\InstallUtil.exe” “C:\MPP\MPP.exe” (See also the „documentation and Interface description“)

## Slide 7: Installation 4/4:
????

## Slide 8: RUN the MPP-Application 1/3:
Without license key*: The application runs only for 2h Start the mpp.exe in the mpp file folder *How to use the license: see „MPP-Configurator“

## Slide 9: RUN the MPP-Application 2/3:
With license key: The application runs as a service Restart PC or Start mpp-service manually if the correct license key is activated

## Slide 10: RUN the MPP-Application 3/3:
With the wrong or without a license key*: The MPP service cannot go into the execution (running) mode. Start the mpp.exe in the mpp folder. The application runs only for 2h. *License key and PrismaPro do not belong together or License key misspelled

## Slide 11: Configuration of MPP 1/4:
Run the MPP-Configurator (Press for more information)

## Slide 12: Configuration of MPP 2/4:
MPP.ini Mainly defines the interfaces such as: IP addresses, license, log files, etc.
OPC_Config.ini Defines the individual OPC tags, their spelling and data types in order to be able to communicate with the OPC-UA server as an OPC-UA client.

## Slide 13: Configuration of MPP 3/4: mpp.ini for S7 and for OPC-UA
S7
OPC-UA The MPP
can only be the Client !!

## Slide 14: Configuration of MPP 4/4: „OPC_Config.ini“ (OPC-UA only)

## Slide 15: MPP .log-Files:
When the application starts, a new file is created with the current time and date and the ending .log or, if the file already exists, the new entries are appended to the end of the file
If no new entries are made, the application runs without any error. Otherwise the entries can be helpful for troubleshooting

## Slide 16: Licensing of MPP: (Without license key the mpp is working for only 2 hours)
The licensing depends on the MAC address of the PrismaPro device
Get the MAC address:
Send this MAC address with your contact details to: QMS@Pfeiffer-Vacuum.de
If you are licensed, you will receive the license key
Fill in the line of the license key in mpp.ini manually or use the MPP-configurator

## Slide 17: Save the configuration:
Copy the Configuration to your MPP-Folder (Where the mpp.exe is)
With a text editor, like notepad, you can edit the ini-Files manually
After a new configuration you have to restart the MPP service or the MPP.exe. (It‘s about with or without license key)

## Slide 18: Application MPP  Siemens PLC 1/3: PLC 1516-3 PN/DP PLC Online (Watch-List)

## Slide 19: Application MPP  Siemens PLC 2/3: PLC 1516-3 PN/DP PLC Online (DB100 Part 1)

## Slide 20: Application MPP  Siemens PLC 3/3: PLC 1516-3 PN/DP PLC Online (DB100 Part 2)

## Slide 21: Example: S7 (direct connection)
*Copyright-Hinweis
Bildquelle "© Siemens AG 2020, Alle Rechte vorbehalten“
192.168.1.100 (Standard IP-Addr.)
X2: 192.168.1.1
X1: 192.168.1.99
192.168.1.101
Siemens PLC
S7-300/400/1500
Siemens TIA
Siemens Box-PC
(MPP.exe + MPP.ini)

## Slide 22: Example: OPC UA Server
192.168.1.100 (Standard IP-Addr.)
X1: 192.168.1.99
192.168.1.98
PC: OPC-UA Server
X1: 192.168.1.199
PC: OPC-UA-Client
Siemens Box-PC
(MPP.exe + MPP.ini + OPC_CONFIG.ini)
*Copyright notice:
Image source"© Siemens AG 2020, All rights reserved

## Slide 23: Example: OPC-UA S7
*Copyright notice:
Image source"© Siemens AG 2020, All rights reserved
192.168.1.100 (Standard IP-Addr.)
X2: 192.168.1.1
X1: 192.168.1.99
192.168.1.101
PC: OPC-UA-Client
S7-PLC: OPC-UA-Server
Siemens Box-PC
(MPP.exe + MPP.ini + OPC_CONFIG.ini)
Siemens PLC

## Slide 24: Example: OPC UA KepwareServerEX6
192.168.1.100 (Standard IP-Addr.)
X1: 192.168.1.99
192.168.1.98
PC: KepwareServerEX 6
X1: 192.168.1.199
PC: OPC-UA-Client
Siemens Box-PC
(MPP.exe + MPP.ini + OPC_CONFIG.ini)
*Copyright notice:
Image source"© Siemens AG 2020, All rights reserved

## Slide 25: Example: OPC UA Rockwell Automation
OPC-UA-Server
FactoryTalk Linx Gateway
192.168.1.100
X2: 192.168.1.99 Ethernet 1
X2 192.168.178.101
X1: 192.168.178.99 Ethernet 2
192.168.178.254
PLC: L30ER
Siemens Box-PC
(MPP.exe + MPP.ini + OPC_CONFIG.ini)
PC: OPC-UA-Client
ROCKWELL
Automation
*Copyright notice:
Image source"© Siemens AG 2020, All rights reserved